# Abstract

Test

# Introduction

[write out full name of DICOM] (DICOM) creates a uniform standard of properties that series of standard properties in imaging which.

DICOM was initially created as a set of standard properties that helps facilitate communication between multiple vendors and technologies in medicine. Modern treatment planning systems (TPS) often contain many features which exist to reduce potential confusion when working with medical images.

In today’s world, there are often circumstances which require changing certain properties of the DICOM images to better enable their use within the clinic. As a commonly seen case in stereotactic radiosurgery, all MRI images (T1, T2, FLAIR) acquired within the same study will have the same frame of reference, or ‘Frame of Reference Unique Identifier (UID)’. This is a feature based on the understanding that the images are acquired in the same location. Unfortunately, this also means that any motion which occurs between scans cannot be corrected, as both images share the same Frame of Reference UID. To ‘break’ this inherent registration, the DICOM value for Frame of Reference UID must be changed on each scan one wishes to register.

Changing any DICOM value can be a difficult task for users who are not overly familiar with the process, and stressful for those who are, with the inherent fear of accidentally changing a different attribute. Certain programs allow for this: MIM [ref] allows the user to anonymize DICOM and change certain values, but this also rewrites many other DICOM attributes, Raystation [ref] allows the used to assign to new frame of reference to an exam, but this is the only attribute that can be changed.

For this reason, we have created the Unlink program. The simple interface offers the option to change three potential values: the Frame of Reference, Series Instance UID, and/or Study Instance UID. Users can specify which modality they would like to change and use the built-in unzip/unzip and run if files need to be extracted before being changed. This program, built in C#, is designed to run on any Windows based computer and is publicly available at [www.GitHub.com/BrianMAnderson/Unzip\_Unlink\_CSharp](http://www.GitHub.com/BrianMAnderson/Unzip_Unlink_CSharp).

# Materials and Methods

The program was written using C#[ref] and .NET framework 4.8, the current .NET standard at time of creation (2023). All DICOM manipulation was facilitated with the FellowOak DICOM package [ref] and SimpleITK[ref]. The main splash screen is shown in Figure 1.




Figure : Main splash screen of the program. There are three check boxes of DICOM attributes that can be changed in the top left and three checkboxes for Modalities to change in the top right.

Users can select any or all the options in the upper left: Frame of Reference, Series Instance UID, and Study Instance UID. Should the

## Installation

The solution can be downloaded directly GitHub, or the pre-built executable can be downloaded and installed from [link to google drive].

# Discussion and Conclusion

Test

# Acknowledgements

Test

# References

Refs