# Introduction

The Digital Imaging and Communications (DICOM1) standard creates a technical protocol for the storage and transmission of medical images and helps facilitate communication between multiple vendors and technologies in medicine.

In radiotherapy clinics, there are often circumstances which require modifying the properties of the DICOM images. As a commonly seen case in stereotactic radiosurgery, all MRI images (T1, T2, FLAIR) acquired within the same study will have the same frame of reference, or ‘Frame of Reference Unique Identifier (UID)’. This is a feature based on the understanding that the images are acquired in the same location. Unfortunately, this also means that any motion which occurs between scans cannot be corrected, as both images share the same Frame of Reference UID. To ‘break’ this inherent registration, the DICOM value for Frame of Reference UID must be changed on each scan one wishes to register.

Changing any DICOM value often requires expert knowledge of file structure or specialized software, posing challenges and potential errors in accidentally or unnecessarily changing other attributes. This is especially challenging when a 4D-CT is manipulated, as the Frame of Refence may need to be changed, but the new value must be consistent across the 4D scan. Barring these changes, the clinic would have to make do with imprecise registrations which add to overall treatment uncertainty.

Modifying DICOM properties is prone to error; attributes can be modified unintentionally, or files can be corrupted and hard to recover2. Commonly used software can modify a subset of DICOM file attributes. For example, MIM3 allows the user to anonymize DICOM and change certain attributes, but this also rewrites many other DICOM attributes such as the date and time of creation. Raystation 10A4 has built-in functionality to assign an exam to a new frame of reference, but this is the only DICOM attribute that can be changed and cannot be performed on examinations with an associated plan.

To address these gaps, we have created the Unlink program which provides a user-friendly interface to change DICOM attributes which are not readily or easily available in commercial software. The simple interface offers the option to change three potential values: the Frame of Reference, Series Instance UID, and/or SOP Instance UID. The SOP Instance UID is defaulted to be always changed, however users should note this will break associations with plan/structure sets.

Users can specify which modality they would like to change and use the built-in unzip feature and run if files need to be extracted before being changed. This feature was added to facilitate an optimized workflow when pulling images from LifeImage which are automatically zipped. This program, built in C#5, is designed to run on any Windows based computer and is publicly available at [https://github.com/<anon>/Unzip\_Unlink\_Csharp](https://github.com/%3canon%3e/Unzip_Unlink_Csharp).

# Methods

The program was tested with publicly available brain MRI scans available here: <https://figshare.com/articles/dataset/Data_from_An_Investigation_of_Machine_Learning_Methods_in_Delta-radiomics_Feature_Analysis/9943334>. This dataset contains several T1 and T2-FLAIR images. Our program was then used to change the series instance UID, frame of reference UID, and study instance UID.

Verification of the edited DICOM was evaluated within the RayStation 10A4 treatment planning system (TPS) and Varian's Eclipse TPS v16.1 (Varian Medical Systems, Palo Alto, CA). Further evaluation was performed with MIM3 to ensure that only the desired attributes were changed in the process.

# Software Format and Usage Notes

The program is written using C# and .NET framework 4.8, the current standard at time of creation (2023). All DICOM manipulation was facilitated with the FellowOak DICOM package6 and SimpleITK7. A demonstration of the program welcome screen is shown in Figure 1.




Figure : Main splash screen of the program. There are three check boxes of DICOM attributes that can be changed in the top left and three checkboxes for Modalities to change in the top right.

Users can select any or all the options in the upper left: Frame of Reference, Series Instance UID, and Study Instance UID, as well as specify which modalities they would like to change: CT, 4DCT, MR, and/or PET images. The specification of modalities is beneficial when multiple modalities are located within the same folder.

Users may wish to change both the Frame of Reference UID and Series Instance UID if multiple registrations are required. For example, when registering a PET/CT to a new simulation scan, it can be useful to register both about the nasal cavity and also about the neck region. Two rigid registrations are sometimes not possible (within RayStation 10A4), and so having two distinct image sets which each has a unique registration can be advantageous to the physician during the delineation process.

New DICOM UIDs are generated using the FellowOak4 C# package.

## 4DCT Registrations

An additional option for 4DCT was added because of the special nature of a 4DCT. Often, a free-breathing scan *and* 4DCT are acquired at the same time. If the user wishes to change the frame of reference UID for the 4DCT, the program will need to create a unique frame of reference UID that is still consistent across all phases of the 4DCT, but distinct from the free-breathing scan.

The program runs in two main steps. First, the program groups all files based on their unique Series Instance UIDs and modalities within the selected folder. Second, the DICOM files associated with each Series Instance UID are loaded, and for each selected attribute (Frame of Reference UID, Series instance UID, Study instance UID) the associated tag is changed with the FellowOak4 package. After all changes have been applied, the new DICOM file is written over the original DICOM file.

When the 4DCT option is selected, any CT with the same frame of reference UID will be given a new frame of reference UID. This can be very useful when a 4DCT has the same frame of reference as a free-breathing scan, and the user wishes to break this registration, but keep the 4DCT together.

## Running the program

*Note that DICOM files are required to be exported from the TPS and need to be in a folder accessible to the user.* Once the DICOM attributes have been changed, the new data can be imported to the TPS.

A visual representation of the entire workflow can be seen in Figure 2. Green bars beneath the ‘Status’ symbol give real-time feedback of the updating process.




Figure : Graphical workflow of program

We noted that, depending on network speed, changing the attributes of a 125 slice CT scan required approximately seven seconds. When the files are located on the local drive there is a significant increase in speed.

### Installation

The solution can be downloaded directly GitHub with pre-built executables or built directly from the source code.

# Results

Publicly available brain MRI images (<https://figshare.com/articles/dataset/Data_from_An_Investigation_of_Machine_Learning_Methods_in_Delta-radiomics_Feature_Analysis/9943334>) and TCIA 4D CT Lung data8 were used as testing images to validate the software. The ability to change the frame of reference, series instance identifier, and study instance identifier using the program was evaluated with both the RayStation 10A4 TPS, MIM3, and Varian’s Eclipse v16.1. Within the TCIA 4DCT Lung data, changing attributes from the native Frame of Reference UID to a new UID that is still consistent across the 4DCT was also evaluated.

# Discussion

The program is designed to run on the Windows operating system and not MAC/Linux. There is concern that institutional internet security division (ISD) may prevent the installation of this program onto a computer. Within our institution we were able to circumnavigate this issue by placing the compiled program on a network drive location which was accessible to the team, which runs without the requirement of installation.

## Potential applications

The program presented here represents an easy, user-friendly method of changing two commonly changed DICOM attributes with a vendor agnostic solution. This can be useful when breaking the inherent registration of consecutively acquired MRI images, registering a free breathing scan to a breath hold scan, or a free breathing scan to a 4D CT. Because we have hosted the tool on GitHub, any user can provide feedback and new attributes can easily be added to the program in the future.

We have implemented this solution within two clinics: <anon> and <anon> with positive feedback from the physics and dosimetry teams. The program is freely available and open for input from the community via GitHub, allowing future updates and improvements as requested.

# Acknowledgements

The authors would like to thank <anon> and <anon> for support of this work.

# Abbreviation List

Treatment Planning System: TPS

# References

1. DICOM. Accessed February 16, 2024. https://www.dicomstandard.org/

2. González DR, Carpenter T, Van Hemert JI, Wardlaw J. An open source toolkit for medical imaging de-identification. *Eur Radiol*. 2010;20(8):1896-1904. doi:10.1007/S00330-010-1745-3/METRICS

3. About MIM Software | MIM Software Inc. Accessed February 15, 2024. https://www.mimsoftware.com/about/mim

4. Bodensteiner D. RayStation: External beam treatment planning system. *Medical Dosimetry*. 2018;43(2):168-176. doi:10.1016/j.meddos.2018.02.013

5. 2013 MC. C# Language Specification Version .NET 4.8.1. Published online 2013. Accessed January 30, 2023. https://dotnet.microsoft.com/en-us/download/dotnet-framework

6. fo-dicom/fo-dicom: Fellow Oak DICOM for .NET, .NET Core, Universal Windows, Android, iOS, Mono and Unity. Accessed July 21, 2022. https://github.com/fo-dicom/fo-dicom

7. I Ã Nez L, Schroeder W, Ng L, Cates J. The ITK Software Guide Second Edition Updated for ITK version 2.4. Published online 2005. Accessed April 18, 2018. http://www.itk.org

8. Hugo GD, Weiss E, Sleeman WC, et al. Data from 4D Lung Imaging of NSC. *The Cancer Imaging Archive*. Published online 2016.

# Figures




Figure 1: Main splash screen of the program. There are three check boxes of DICOM attributes that can be changed in the top left and three checkboxes for Modalities to change in the top right.




Figure 2: Graphical workflow of program